SIO Shipboard Meteorological Sensor Specifications

1. General information	2

2.1 Underway water measurement sensors used on SIO vessels	3

3. R/V New Horizon	4

3.1 Sensors located on forward part of the upper bridge	4

3.2 Sensors located in Aft Lab:	5

3.3 Hull Mounted Sensors	5

4. R/V Sproul	6

4.1 Sensors located on top bridge deck	6

5. R/V Revelle	7

5.1 Sensors located on Forward MET Mast:	7

5.2 Sensors located in Aft Hydro LAB:	8

5.3 Sensors located in Bow Thruster Room	8

6.  R/V Melville	9

6.1 Sensors located on Forward MET Mast	9

6.2 Sensors located in Bio/Analytical Lab	10

6.3 Sensors located in Bow Chamber	10

Rev 1.15 14 May 2003

# 1. General information

SIO vessels are usually outfitted with the following Meteorological sensors:

Air Temperature

Barometric Pressure

Relative Humidity

Wind Speed/Direction

Long Wave Radiation

Short Wave Radiation

Precipitation

Sea Surface Temperature

*Thermosalinograph

On R/V Revelle and R/V Melville additional sensors located in the lab measure underway

oxygen and fluorometer.

The meteorological sensors used on SIO vessels are unmodified from the original manufacturers specifications with the exception of the installation of a common type of connector. This allows the sensors to be interchangeable between all SIO ships so that the sensors can be easily replaced and periodically rotated between ship and shore for maintenance/calibrations.

Depending on the make and model of the sensor the data output may be RS485, RS232 or analog signal out.

The mast junction box that is installed on each ship is designed to accommodate sensors of different output types. There are also provisions for expansion to support specific science expedition requirements.

MET Sensors are calibrated at a 6-12 months periodicity. Sensor accuracy and calibration specifications as listed are valid if the sensors are in good working order and have been calibrated within the past six months.

All sensor tolerance and accuracy values listed in this document are taken from the manufacturers specification sheets. They are derived under laboratory conditions and do not necessarily indicate the accuracies that may be expected in actual shipboard installations. They do not take into account the effects of ship roll, pitch, sway, solar radiation, shipboard structural interference and thermal radiation effects of the ship itself.  These effects have not been determined and will vary from ship to ship. Placing the sensors forward on the ship and as high as possible minimizes some of the ship structural radiation effects. This is the case on R/V Revelle and R/V Melville, which has a bow, mounted MET mast.

* Installed upon request prior to cruise (R/V New Horizon, R/V Sproul).

2. Meteorological Sensors used on SIO vessels

| Type | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Air Temperature | RM Young | 41342LC | -50.0 to +50.0 C | +/- 0.3 Deg C | 4-20 ma |
| Precipitation | RM Young | 50202 | 0-50 mm | +/- 1.0 mm | 0 - 5 Volts |
| Wind | RM Young | 05103 | Dir 0-360 Deg Spd 0-60 m/s | +/- 3.0 Deg +/- 0.3 m/s | Voltage Out Pulse Count |
| Humidity/Temp | Rotronics | MP101A | RH 0-100% Tmp –40 to +60 DegC | +/- 1.5% +/- 0.2 Deg C | 0 – 1 VDC -0.4 - 0.6 VDC |
| Humidity/Temp | Vaisala | HMP45A | RH 0-100% Tmp –40 to +60 C | +/-1.0% +/- 0.2 Deg C | 0 – 1 VDC 0 – 1 VDC |
| Barometer | Air Inc. | DB-2A | 800-1060mb | +/- 0.5 mb | RS232 |
| Barometer | Vaisala | PTB100A | 800-1060mb | +/- 0.3 mb | 0 – 5.0VDC |
| Barometer | Vaisala | PTB101C | 900-1100 mb | +/- 0.3 mb | 0 - 2.5VDC |
| Radiometer Long Wave | Eppley | PIR Pyrgeometer | 3.5-50 um 4 uv/W/M2 | +/- 1% Linearity | 0 – 10mv |
| Radiometer Short Wave | Eppley | PSP Pyranometer | 285-2800nm 9 uv/W/M2 | +/- 0.5% Linearity | 0 – 100mv |
| Radiometer Black&White | Eppley | 8-48 | 285-2800nm 10 uv/W/M2 | +/- 1% Linearity | 0 – 100mv |
| Surface PAR | Biospherical | QSR-240P | 400-700nm 1.4e-5uE/(cm2-sec) to 0.5uE/(cm2-sec) | Not Specified | 0 - 5VDC |

## 2.1 Underway water measurement sensors used on SIO vessels

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Thermosalinograph | Seabird | SBE-21 | Temp –5 to +35 C Cond 0 to 65 mS/cm | +/- 0.01 Deg C +/- 0.01 mS/cm | RS232 |
| Thermosalinograph | Seabird | SBE-45 | Temp –5 to +35 C Cond 0 to 70 mS/cm | +/- 0.002 Deg C +/- 0.003 mS/cm | RS232 |
| Thermosalinograph | FSI | OTM OCM | Temp –2 to +32 C Cond 0 to 70 mS/cm | +/- 0.003 Deg C +/- 0.005 mS/cm | RS485 |
| Fluorometer | Wetlabs | WetStar | 0.03 to 75 ug/l | Not Specified | 0 - 5VDC |
| Water Temperature | Omega | ON-403-PP | Temp –5 to +35 C | +/- 0.02 Deg C | 0 - 5VDC |
| Oxygen | Seabird | SBE-43 | 120% surf saturation | 2% | 0 - 5VDC |
| Hull mounted Sea Surf Temp | STS | SEG-14 | Temp –2 to +35 C | +/- 0.1 Deg C | RS232 |
| Flow Meter | Signet | 515-PO | 0-80 GPM | +/- 1.0% FS | Pulse Count |
| Flow Meter | FLO-CAT | C-ES45-B002 | 0.27 – 18.9 LPM | +/- 1.0% FS | Pulse Count |
| Flow Meter | FLO-CAT | C-ES45-B003 | 0.38 – 37.9 LPM | +/- 1.0% FS | Pulse Count |

# 3. R/V New Horizon

Current MET configuration:

Meteorological Sensors on Top bridge forward rail

Thermosalinograph in aft lab

Sea Surface Temperature sensors (Hull mounted)

GPS Pcode

Gyro

Knudsen Water Depth

Distance from Mean Water Line to sensors - About 36'

## 3.1 Sensors located on forward part of the upper bridge

| Type | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Air Temperature | RM Young | 41342LC | -50.0 to +50.0 C | +/- 0.3 Deg C | 4-20 ma |
| Wind | RM Young | 05103 | Dir 0-360 Deg Spd 0-60 m/s | +/- 3.0 Deg +/- 0.3 m/s | Voltage Out Pulse Count |
| Humidity Temperature | **Vaisala | HMP45A | RH 0-100% Temp –40 to +60 C | +/-1.0% +/- 0.2 Deg C | 0 – 1VDC 0 – 1 VDC |
| Barometer | Vaisala | PTB101C | 900-1100 mb | +/- 0.3 mb | 0 – 2.5VDC |
| Radiometer Long Wave | Eppley | PIR | 3.5-50 um 4 uv/W/M2 | +/- 1% Linearity | 0 – 10mv |
| Radiometer Short Wave | Eppley | PSP | 285-2800nm 9 uv/W/M2 | +/- 0.5% Linearity | 0 – 100mv |
| * Surface PAR | Biospherical | QSR-240P | 400-700nm 1.4e-5uE/(cm2-sec) to 0.5uE/(cm2-sec) | Not Specified | 0 - 5VDC |

All above sensors plug into a MET Sensor junction box that utilizes analog to RS485 converters as well as RS232 to RS485 converters.

The Mast Junction box is SEG Type 8B and can accommodate the following number of sensor types.

1 RM Young Anemometer connection

2 RS485 connections

1 RS232 input

7 Analog voltage inputs

2 Analog 4-20ma inputs

Air Temperature and Humidity sensors are installed in a multi-plate radiation shield (RM Young 41002).

The Barometer utilizes a static pressure head – Vaisala SPH-10

** The Humidity/temperature sensor is typically the Vaisala HMP-45A. On occasion the Rotronics MP101A may replace it.

## 3.2 Sensors located in Aft Lab:

| TYPE | MFG | Model | Range | Accuracy | OutPut |
| --- | --- | --- | --- | --- | --- |
| *Thermosalinograph | Seabird | SBE-21 | Temp –5 to +35 Deg C Cond 0 to 65 mS/cm | +/- 0.01 Deg C +/- 0.01 mS/cm | RS232 |
| * Fluorometer | Wetlabs | WetStar | 0.03 to 75 ug/l | Not Specified | 0 – 5VDC |

The Fluorometer signal is connected to one of the 4 channels A/D built into the SBE-21. The SBE-21 RS232 output is connected to the main MET Acquisition computer.

## 3.3 Hull Mounted Sensors

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Hull mounted Sea Surface Temp | STS | SEG-14 | Temp –2 to +35 Deg C | +/- 0.1 DegC | RS485 |

These sensors are located in the machine shop compartment, starboard side, between the hull and insulation.

The RS485 Output is connected to the main MET Acquisition computer.

* Installed upon request prior to cruise

# 4. R/V Sproul

Current MET configuration:

Meteorological Sensors on Flying bridge

GPS Pcode

Gyro

Knudsen Water Depth

A description of the sensors that are on the R/V Sproul follows:

Coastal Environmental WeatherPak, Sample rate once every 10 seconds

Location: Top bridge Deck

Distance Sensors Mounted above Top Bridge Deck - about 7'

Distance from Mean Water Line to sensors - About 33'

## 4.1 Sensors located on top bridge deck

| Type | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Weather station | Coastal Environmental | Weatherpak |  |  | RS232 |
| Wind | RM Young | 05103 | Dir 0-360 Deg Spd 0-60 m/s | +/- 3.0 Deg +/- 0.3 m/s | Weatherpak |
| Barometer | Coastal Environmental | 1081 | 300-1100mb | +/- 1mb | Weatherpak |
| Humidity/Temp | Vaisala | HMP45A | RH 0-100% Temp –40 to +60 C | +/-1.0% +/- 0.2 Deg C | Weatherpak |
| Radiometer Long Wave | Eppley | PIR | 3.5-50 um 4 uv/W/M2 | +/- 1% Linearity | Weatherpak |
| Radiometer Black&White | Eppley | 8-48 | 285-2800nm 10 uv/W/M2 | +/- 1% Linearity | Weatherpak |

All above sensors plug into the Weatherpak. The RS232 output of the weatherpak feeds the main MET acquisition computer located in the upper Lab van.

Air Temperature and Humidity sensors are installed in a multi-plate radiation shield (RM Young 41002).

# 5. R/V Revelle

Meteorological Sensors on Forward MET Mast

Thermosalinograph Sensors in Bow Thruster Room

Flo-thru Sensors in Aft Hydro lab

Wind Sensors on Main Mast

Timeserver for precision Time Stamping of data

GPS Pcode for Ships position

Sperry MK37 Gyro or TSS Meridian Surveyor for ships heading

Knudsen Water Depth

MET system outputs:

Sound Velocity to EM120 Simrad Multibeam system

The following Sensors are mounted on the Forward MET mast.

There are two levels.

- First Platform - 56’above Mean Water Line.

- Second Platform - 68’ above Mean Water Line

## 5.1 Sensors located on Forward MET Mast:

| Type | MFG | Model | Range | Accuracy | Output | MWL |
| --- | --- | --- | --- | --- | --- | --- |
| Air Temp | RM Young | 41342LC | -50.0 to +50.0 C | +/- 0.3 Deg C | 4-20 ma | 57’ |
| Precipitation | RM Young | 50202 | 0-50 mm | +/- 1.0 mm | 0 - 5 Volts | 57’ |
| Wind | RM Young | 05103 | Dir 0-360 Deg Spd 0-60 m/s | +/- 3.0 Deg +/- 0.3 m/s | Volts Pulse Count | 59’ |
| Humidity Temperature | **Vaisala | HMP45A | RH 0-100% Tmp –40 to +60 C | +/-1.0% +/- 0.2 Deg C | 0 – 1 VDC 0 – 1 VDC | 57’ |
| Barometer | **Air Inc. | DB-2A | 800-1060mb | +/- 0.5 mb | RS232 | 57’ |
| Radiometer Long Wave | Eppley | PIR | 3.5-50 um 4 uv/W/M2 | +/- 1% Linearity | 0 – 10mv | 68’ |
| Radiometer Short Wave | Eppley | PSP | 285-2800nm 9 uv/W/M2 | +/- 0.5% Linearity | 0 – 100mv | 68’ |
| * Surface PAR | Biospherical | QSR-240P | 400-700nm 1.4e-5uE/(cm2-sec) to 0.5uE/(cm2-sec) | Not Specified | 0 - 5VDC | 68’ |

All above sensors plug into a MET Sensor junction box that utilizes analog to RS485 converters as well as RS232 to RS485 converters. A four-conductor cable coming from the Bos’n locker connects the RS485 feed to the SIS cable system, which is then connected to the main MET acquisition computer.

The Mast Junction box is SEG Type 8A and can accommodate the following number of sensor types.

1 RM Young Anemometer connection

4 RS485 connections, 2 RS232 inputs

7 Analog voltage inputs, 4 Analog 4-20ma inputs

Air Temperature and Humidity sensors are installed in a multi-plate radiation shield (RM Young 41002). The Barometer utilizes a static pressure head – Vaisala SPH-10

** The Humidity/temperature sensor is typically the Vaisala HMP-45A. On occasion the Rotronics MP101A may replace it. The barometer is typically the AIR Inc DB-2A but on some occasions the Vaisala PTB-100A or PTB-101C may replace it.

* Installed upon request prior to cruise

## 5.2 Sensors located in Aft Hydro LAB:

Uncontaminated Seawater from bow feeds a vortex debubbler. The debubbled water is then fed into the flo-thru system.

Flo-thru System

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Fluorometer | Wetlabs | WetStar | 0.03 to 75 ug/l | Not Specified | 0 - 5VDC |
| Water Temperature | Omega | ON-403-PP | Temp –5 to +35 C | +/- 0.02 Deg C | 0 - 5VDC |
| Oxygen | Seabird | SBE-43 | 120% surf saturation | 2% | 0 - 5VDC |
| FlowMeter | FLO-CAT | C-ES45-B003 | 0.38 – 37.9 LPM | +/- 1.0 % FS | Pulse Count |
| * Thermosalinograph | Seabird | SBE-45 | Temp –5 to +35 C Cond 0 to 70 mS/cm | +/- 0.002 Deg C +/- 0.003 mS/cm | RS232 |

The above sensors plug into a Sensor junction box that utilizes analog to RS485 converters. The RS485 feed connects via the SIS system to the Main MET acquisition computer in computer Lab.

## 5.3 Sensors located in Bow Thruster Room

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Thermosalinograph | Seabird | SBE-45 | Temp –5 to +35 C Cond 0 to 70 mS/cm | +/- 0.002 Deg C +/- 0.003 mS/cm | RS232 |
| Flow Meter | Flocat | C-ES45-B002 | 0.27 – 18.9 LPM | +/- 1.0% FS | Pulse Count |

The Thermosalinograph is located about 7’ from the seawater intake.

An LCD panel provides a local readout for Sea Surface Temperature, Salinity and water flow in the uncontaminated seawater supply.

* Installed upon request prior to cruise

# 6.  R/V Melville

Current MET configuration:

Meteorological Sensors on Forward MET Mast

Thermosalinograph Sensors in Bow Chamber

Flo-thru Sensors in Analytical Lab

Timeserver for precision Time Stamping of data

GPS Pcode

Sperry MK37 Gyro

Knudsen Depth

MET system outputs:

Navigation string to Seabeam 2000

RS232 data out to main lab splitter

The following Sensors are mounted on the forward mast at a location 55’ above the mean water line:

## 6.1 Sensors located on Forward MET Mast

| Type | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Air Temp | RM Young | 41342LC | -50.0 to +50.0 C | +/- 0.3 Deg C | 4-20 ma |
| Precipitation | RM Young | 50202 | 0-50 mm | +/- 1.0 mm | 0 - 5 Volts |
| Wind | RM Young | 05103 | Dir 0-360 Deg Spd 0-60 m/s | +/- 3.0 Deg +/- 0.3 m/s | Voltage Out Pulse Count |
| Humidity/Temp | Vaisala | HMP45A | RH 0-100% Tmp –40 to +60 C | +/-1.0% +/- 0.2 Deg C | 0 – 1 VDC 0 – 1 VDC |
| Barometer | Vaisala | PTB101C | 900-1100 mb | +/- 0.3 mb | 0 - 2.5VDC |
| Radiometer Long Wave | Eppley | PIR | 3.5-50 um 4 uv/W/M2 | +/- 1% Linearity | 0 – 10mv |
| Radiometer Short Wave | Eppley | PSP | 285-2800nm 9 uv/W/M2 | +/- 0.5% Linearity | 0 – 100mv |
| * Surface PAR | Biospherical | QSR-240P | 400-700nm 1.4e-5uE/(cm2-sec) to 0.5uE/(cm2-sec) | Not Specified | 0 - 5VDC |

All above sensors plug into a MET Sensor junction box that utilizes analog to RS485 converters as well as RS232 to RS485 converters. A four-conductor cable coming from the Bos’n locker connects the RS485 feed to the SIS cable system, which is then connected to the main MET acquisition computer.

The Mast Junction box is SEG Type 8B and can accommodate the following number of sensor types.

1 RM Young Anemometer connection

2 RS485 connections, 1 RS232 input

7 Analog voltage inputs, 2 Analog 4-20ma inputs

Air Temperature and Humidity sensors are installed in a multi-plate radiation shield (RM Young 41002). The Barometer utilizes a static pressure head – Vaisala SPH-10

** The Humidity/temperature sensor is typically the Vaisala HMP-45A. On occasion the Rotronics MP101A may replace it. The barometer is typically the Vaisala PTB-101C but on some occasions it may be replaced by the AIR Inc. DB-2A.

* Installed upon request prior to cruise

## 6.2 Sensors located in Bio/Analytical Lab

Flo-thru System

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Fluorometer | Wetlabs | WetStar | 0.03 to 75 ug/l | Not Specified | 0 - 5VDC |
| Water Temperature | Omega | ON-403-PP | Temp –5 to +35 C | +/- 0.02 Deg C | 0 - 5VDC |
| Oxygen | Seabird | SBE-43 | 120% surf saturation | 2% | 0 - 5VDC |
| FlowMeter | FLO-CAT | C-ES45-B003 | 0.38 – 37.9 LPM | +/- 1.0% FS | Pulse Count |
| * Thermosalinograph | Seabird | SBE-45 | Temp –5 to +35 C Cond 0 to 70 mS/cm | +/- 0.002 Deg C +/- 0.003 mS/cm | RS232 |

The above sensors plug into a Sensor junction box that utilizes analog to RS485 converters. The RS485 feed connects to the Main MET acquisition computer in computer Lab.

## 6.3 Sensors located in Bow Chamber

| TYPE | MFG | Model | Range | Accuracy | Output |
| --- | --- | --- | --- | --- | --- |
| Thermosalinograph | FSI | OTM OCM | Temp –2 to +32 C Cond 0 to 70 mS/cm | +/- 0.003 Deg C +/- 0.003 mS/cm | RS485 |
| Flow Meter | Signet | 515-PO | 0-80GPM | +/- 1.0 % FS | Pulse Count |

The Thermosalinograph is located about 4’ from the seawater intake.

The above sensors plug into a Sensor junction box. The RS485 feed connects to the Main MET acquisition computer in computer Lab.

* Installed upon request prior to cruise